Contents lists available at ScienceDirect

Radiation Physics and Chemistry  journal homepage: www.elsevier.com/locate/radphyschem

| # New developments in monochromatic energy dispersive X-ray fluorescence instrumentation for monitoring toxic elements in food matrices: Advantages and limitations Boris Johnson-Restrepo a, Ezekiel Blain b,c, Christopher Judd b,c, Andrew Tysoe b, Patrick J. Parsons b,c,* a Environmental Chemistry Research Group, School of Exact and Natural Sciences, San Pablo University Campus, University of Cartagena, 130015, Cartagena, Colombia b Division of Environmental Health Sciences, Wadsworth Center, New York State Department of Health, USA c Department of Environmental Health Sciences, College of Integrated Health Sciences, University at Albany, Rensselaer, USA |  |  |
| --- | --- | --- |
| A R T I C L E I N F O |  | A B S T R A C T |
| Handling Editor: Dr. Chris Chantler |  | Monochromatic excitation (M-) based on a low-power X-ray (20–50W) source with doubly curved crystal (DCC) optics has been used successfully for X-ray fluorescence (XRF) analysis of many different sample matrices. Previously, a“first-generation”M-XRF analyzer, the HD Mobile® (X-ray Optical Systems, (XOS) East Greenbush NY), was used in a field-based public health study. The HD Mobile® was optimized to detect Pb, Cd, Hg and As in consumer products, food, and cultural medications, with limits of detection (LOD) on the order of a single digit μg/g (ppm) mass fraction. In this current study, two“second generation”DCC-enabled XRF analyzers were evaluated for detecting key toxic elements in food matrices. The two instruments evaluated included: the Zmax™ XRF Analyzer and the Emax™ Analyzer, which are both manufactured by Z-Spec (East Greenbush, NY, US). The Zmax™ uses an X-ray tube operated at 35W with a single DCC optic for a 17.5-keV excitation beam optimized to detect Pb, As, Hg, Ni and Cr. A fast silicon drift detector (SDD) is used to detect fluorescence x-rays. The Emax™ uses an X-ray tube operated at 40 W with DCC optics for a 30-keV excitation beam, and a thicker 1-mm SDD optimized to detect Cd, Mo, Sb, and Sn, along with many other elements. While the primary focus of the current study was the Emax™ and Zmax™, several other XRF analyzers were available and were included for com- parison purposes. They included the HD Mobile® (XOS), and two ζconventional’portable XRF analyzers: the Niton™ XL3t GOLDD XRF, and the Niton™ XL5 Plus, both manufactured by ThermoFisher Scientific (Tewksbury, MA). Z-Spec used certified reference materials (CRMs) to optimize the Fundamental Parameters algorithm for the Zmax™ and Emax™ analyzers to achieve acceptable performance for key toxic elements in biological matrices. For the Zmax™, difference plots show good agreement between found values and CRM assigned values for the elements studied (Pb, As, Cu, Mn, Ni, Hg and Zn). For the Emax™, there was good agreement for values obtained for Cd in CRMs. Archived food-based proficiency test (PT) samples were analyzed to assess the performance of the Zmax™ and Emax™ as screening tools to detect key elements. Some of the PT samples were also analyzed using the two Niton™ XRF instruments as well as the HD Mobile® for comparison purposes. These PT samples were spiked with toxic elements and circulated by the FDA to US-based public health laboratories funded under the Food Emergency Response Network (FERN), with target values assigned based on consensus data obtained from analyses based on Inductively Coupled Plasma Mass Spectrometry (ICP-MS). For the FERN 2022 PT samples (sausage meat), which were spiked with Tl, Cd and Pb, Zmax™ values were within the tolerance range required for ICP-MS analysis. Estimated LODs in μg/g (ppm) were: Pb, 0.06; Tl, 0.06; Hg, 0.06; Cd, 0.04; and Sn, 0.11. Several ζreal-world’samples were analyzed using the Zmax™ and Emax™ devices, including medicinal herbs such as moringa and pot marigold, as well as various spices like cloves, coriander, and curry. |
| Keywords: Monochromatic XRF Toxic metals Food Validation |  |  |

This article is part of a special issue entitled: ISRP-16 published in Radiation Physics and Chemistry.

*  Corresponding author. Division of Environmental Health Sciences, Wadsworth Center, New York State Department of Health, USA. E-mail address: patrick.parsons@health.ny.gov (P.J. Parsons).

https://doi.org/10.1016/j.radphyschem.2025.112749

Received 14 January 2025; Received in revised form 19 March 2025; Accepted 23 March 2025 Available online 24 March 2025

0969-806X/© 2025 Elsevier Ltd. All rights are reserved, including those for text and data mining, AI training, and similar technologies.

## 1.  Introduction

X-ray fluorescence (XRF) is a well-established technique for deter- mining  the  elemental  composition  of a  sample by measuring  the  in- tensity  of X-rays  emitted when  the  sample  is  subjected to  a  suitable excitation source. This non-destructive technique is effective for both qualitative and quantitative analysis and is used extensively in many different  industries  including  petrochemical,  metallurgical  for  metal finishing  and refining,  food,  plastics  and  polymers, textiles,  pharma- ceuticals, cosmetics, forensics, wood treatment, art, antiques and arti- facts    analysis,    chemical    process    control,    and    for    monitoring environmental contaminants (Marguí et al., 2022; Revenko and Pash- kova, 2023; Singh et al., 2022). In general, XRF spectrometers used for analytical  applications  typically  follow  two  approaches:  wavelength dispersive (WDXRF) and energy dispersive (EDXRF) analysis. WDXRF employs a crystal optic to select a specific X-ray wavelength, providing high  resolution  for  elemental  specificity  and  sensitivity.  In  contrast, EDXRF   detects   multiple   elements   simultaneously,   with   resolution contingent   on    detector    performance.    While    WDXRF    excels   in signal-to-background ratio, EDXRF offers broader elemental coverage. Each  XRF  approach  has  its  distinct  strengths  in  elemental  analysis, balancing specificity and range.

Monochromatic excitation (M-) XRF is a particular implementation of EDXRF  that  employs  a  single-wavelength  X-ray  beam  to  enhance element detection by minimizing interferences from other elements and background noise (Chen et al., 2008; Chen and Wittry, 1998). Portable M-XRF  instruments  utilize  a  low-power  X-ray  source  (20–50  W)  in conjunction with a doubly curved crystal (DCC) optic, which is essential for generating a monochromatic X-ray beam (Chen and Gibson, 2002). The DCC serves as a monochromator for two primary purposes. First, it enhances the intensity of diffracted beams from point X-ray sources by optimizing  the  Bragg  condition  at  a  wider  range  of incident  angles. Finally, it enables concentration and focusing of the output X-ray beam. These  benefits  arise  from  the  uniform  curvature  of the  plate,  which significantly enhances its functionality in X-ray diffraction applications (Chen and Gibson, 2002). This instrumentation setup enables detection of elements at very low concentrations with good precision and accu- racy, making it a key analytical technique for increased sensitivity and specificity. M-XRF also offers improved detection limits by minimizing spectral interferences, enabling the identification of trace elements, and by boosting the signal-to-background ratio. This capability is essential for monitoring low concentrations of toxic elements quickly and easily across various matrices, establishing it as a vital tool in environmental investigations (Guimar et al., 2016). Effective public health moni- toring  requires  rapid  detection  of toxic  elements  in  diverse  samples, including   food,   clinical,   environmental,   and   consumer   products. Field-based and lab-based investigations benefit from XRF technologies that offer both portability and high analytical performance (Kalnicky and  Singhvi,  2001).  One  of the  first  commercial  systems  to  include M-XRF for field-based studies was the HD-Mobile® (X-ray Optical Sys- tems, East Greenbush, NY). The HD Mobile® was designed to accom- modate  a  low-power  X-ray  source  as  well  as  a  means  for  quickly switching only one DCC optic that provides efficient signal collection for elements  such  as Pb,  Hg  and As down to  single-digit  μg/g  detection limits  (Fleming et al.,  2017, 2020; Guimar et al.,  2016; McIntosh et al., 2016). The analyzers Zmax™ and Emax™ currently manufactured by Z-Spec Inc. (East Greenbush, NY) represent the next-generation of M-XRF devices that promise even lower detection limits for some key toxic elements such as Cd and Pb.

Medicinal  herbs  and  spices  are  commonly  derived  from  various plants  components,  including fruits,  seeds, barks, flowers, roots,  and rhizomes. Many different plants have been utilized for millennia for this purpose. Beyond their use as food flavorings, some spices also play a significant role in traditional medicine systems across different cultures, such as Ayurveda, Traditional Chinese Medicine, and European herbal practices (Acheampong et al., 2024; Akhbarizadeh et al., 2023). Despite

their  numerous  benefits,  spices,  like  all  plant-based  products,  can become contaminated with toxic metals (and metalloids) during culti- vation, harvesting, and post-harvest processing stages. Chronic exposure to even low levels of toxic metals such as Cd, Pb, Hg, and the metalloid As, can result in bioaccumulation within the body and have been asso- ciated with a range of adverse health issues, including kidney disease, neurological disorders, and developmental delays in children (Ahsan, 2011).  Since  spices  are  often  used  in  small  amounts  but  consumed frequently, even trace levels of contamination can have long-term health implications. Considering these risks, there is a growing awareness to improve  monitoring  and  regulation  of  spices  and  herbal  medicinal products. International organizations such as the World Health Orga- nization (WHO) and the Food and Agriculture Organization (FAO) have set guidelines for maximum allowable levels of contaminants in food products, including spices (WHO, 2007).

To  monitor  toxic  metals  in  spices  and  medicinal  herbs,  reliable analytical techniques are required. Traditional high complexity methods such as atomic absorption spectrometry (AAS) and inductively coupled plasma mass spectrometry (ICP-MS) are widely used for detecting and quantifying trace elements in complex matrices. These techniques pro- vide high sensitivity and precision, but they are also time-consuming and  require  expensive  equipment  and  specialized  personnel.  Recent advances  with  XRF  instrumentation,  offers  an  alternative  screening approach that is rapid and non-destructive, with lower operating costs and requires less skill to operate compared to AAS and ICP-MS.

The principal goals of this study were: (1) to evaluate the perfor- mance  of  two  second-generation  M-XRF  analyzers,  the  Zmax™  and Emax™ devices (Z-Spec, Inc., East Greenbush, NY), for a suite of key toxic elements in food matrices; (2) to compare the Zmax™ and Emax™ with the HD-Mobile® (X-ray Optical Systems, East Greenbush, NY), a first-generation M-XRF analyzer; (3) to compare the M-XRF analyzers to the Niton™ XL3t GOLDD and Niton™ XL5 Plus analyzers (ThermoFisher Scientific, Tewksbury, MA), two conventional polychromatic XRF ana- lyzers; and (4) to analyze a selection of real world medicinal herbs and spices for a suite of toxic elements including Hg, As, Pb, and Cd, as wells as Fe, Mn and Zn using the Zmax™ and Emax™ instruments.

## 2.  Materials and methods

### 2.1.  XRF instrumentation

#### 2.1.1.  Z-spec M-XRF analyzers

Measurements were performed using the Zmax™  and Emax™  M- XRF  analyzers  (Z-Spec,  East  Greenbush,  NY).  These  XRF  analyzers (Fig.  1)  are  equipped with  DCC  optic(s)  for  M-XRF  excitation  (Chen et al., 2008; Chen and Wittry, 1998). The M-XRF instrumentation em- ploys a low power polychromatic X-ray source that is then focused to a three-dimensional  monochromatic  X-ray  beam  by  a  DCC  optic.  This setup allows for the simultaneous excitation of all the elements of in- terest  in  the  sample. A  fast  silicon-drift  detector  (SDD)  captures  the fluorescence X-rays for processing into an energy dispersive spectrum. The Zmax™ uses an X-ray tube operated at 35W and a single DCC optic for an approximate 17-keV excitation beam optimized to detect Pb, As, Hg, Ni and Cr. The Emax™ uses an X-ray tube operated at 40 W with DCC  optics  for  a  30-keV  excitation  beam,  and  a  thicker  1-mm  SDD optimized to detect Cd, Mo, Sb, and Sn. The beam focus for both the Zmax™ and Emax™ is approximately 2 mm × 1 mm.

Both instruments utilize a Fundamental Parameters (FP) model to obtain elemental concentrations from the spectrum. While use of FP is well established for XRF (de Boer,1989; Kitov, 2000) each manufacturer uses a proprietary algorithm to solve the FP equations. Briefly, the raw spectrum is processed by a spectral fitting algorithm to obtain the net count rate for each the elements detected in the sample. Then these net intensities are used to solve the FP equations to obtain elemental con- centrations.  The  Z-Spec  Zmax™  instrument  is  designed  to  achieve optimal Limits of Detection (LOD)s for the elements Pb, Hg, As, Se, and

Fig. 1. Monochromatic-excitation XRF (M-XRF) showing the general optical schematics, for (a) Z-Spec Emax™ and (b) Z-Spec Zmax™ Analyzers.

Tl. In contrast, the Emax™ is designed to achieve optimal LODs for Mo, Cd, Sn, and Sb. Nonetheless, the Zmax™ has the capability to report data for Pb, As, Hg, Tl, Bi, As, Se, Br, Sr, Gd, Ge, Zn, Cu, Ni, Co, Fe, Mn, Ti, Cr, Ca, K, and Cl. In contrast, the Emax™ has the capability to report data for Cd, Sn, Sb, Ag, Mo, Y, Zr, Sr, Br, I, Rb, Pb,As, Hg, Zn, Cu, and Mn.

#### 2.1.2.  X-ray Optical Systems (XOS) HD-Mobile®

Measurements  were  also  performed  using  an  HD  Mobile®  XRF analyzer (X-ray Optical Systems, East Greenbush, NY), a first-generation M-XRF device. The XOS HD Mobile® is equipped with a Mo-anode X-ray excitation source, which is operated at 50 kV and 0.1 mA current. The XRF analyzer incorporates three DCC optics that focus the primary X-ray beam to ~1-mm spot. Complete details of the HD Mobile®ζs operating system and performance characteristics have been reported elsewhere (Guimar et al., 2016). Previous work with the HD Mobile® reported that the“plastic”measurement mode was the most suitable for biolog- ical matrices and, like other XRF manufacturers, uses a proprietary FP software program (i.e., the XOS“Solver”) for real time data processing (Guimar et al., 2016; McIntosh et al., 2016). Although some of the LODs previously reported for the HD Mobile® are limited to single digit ppm (μg/g), it seemed reasonable to assume that some toxic elements supplemented into food matrices might yield detectable values. There- fore,  data  from  the  HD  Mobile®  are  included  here  for  comparison purposes.

#### 2.1.3.  Thermo Scientific Niton™ XRF analyzers

Two conventional Thermo Fisher Niton™ XRF analyzers were also available  for  this  study,  and  data  are  included  here  for  comparison purposes,  where  appropriate.  An  older  Niton™  XL3t  GOLDD  XRF analyzer and a newer Niton™ XL5 Plus XRF analyzer (Thermo Scientific, Tewksbury, MA). The Niton™ XL3t utilizes a 50 kV, 40 μA Au tube to generate a 2W excitation beam, and a Geometrically Optimized Large Area  Drift  Detector  (GOLDD).  It  employs  a  proprietary  FP  model  to determine the mass fraction of the sample for various materials in μg/g (ppm). For the analysis of Proficiency Test (PT) food samples, the de- tector was run in“plastics”calibration mode as that offered the most similar  detection  matrix.  An  optimal  2-min  measurement  time  was suggested by the manufacturer to allow for the lowest possible LODs while  limiting  the  overheating  of  the  detector.  The  detector  was mounted in a Thermo Scientific SmartStand which allowed for better repeatability of results since the detector did not move between mea- surements. The Niton™ XL5 Plus uses  an Ag tube to  generate  a  5W excitation beam and employs a 6-position filter for enhanced spectral coverage. The beam is collimated to an 8-mm spot size which is incident on a custom extra-large area SDD. Like the XL3t, the XL5 Plus mea- surements were performed in“plastics”calibration mode with a 2-min

run  time.  The  detector  was  housed  in  a  new  Thermo  Portable  Test Stand which is like the SmartStand and is designed to prevent the de- tector   from   moving   between   measurements   allowing   for   better repeatability.

### 2.2.  Sample materials for XRF analysis

#### 2.2.1.  Certified reference materials

Various certified reference materials (CRMs) were analyzed using the two Z-Spec M-XRF analyzers to evaluate accuracy (where possible). The CRMs  selected  for  analysis  were  procured  from  different  National Metrological Institutes (NMIs): a) National Institute of Standards and Technology (NIST, Gaithersburg, MD); NIST Standard Reference Mate- rial (SRM) 1515 Apple Leaves; NIST SRM 1577c Bovine Liver Tissue; NIST SRM 1547 Peach Leaves; NIST SRM 1568b Rice Flour; NIST SRM 1573a  Tomato  Leaves;  NIST  SRM  1570a  Spinach  Leaves;  NIST  SRM 2976 Muscle Tissue (trace elements and methylmercury); and NIST SRM 2977  Muscle  Tissue  (Organic  Contaminants  and  Trace  Elements);  b) National  Research  Council  of  Canada  (NRC,  Ottawa,  Canada),  NRC BOVM-1 Bovine Muscle Certified Reference Material for Trace Metals and other Constituents; (c) International Atomic Energy Agency (IAEA, Vienna, Austria), IAEA-413 Major, Minor, And Trace Elements in Algae; (d) European Commission, Joint Research Centre (ERM, Geel, Belgium), ERM-BD151 Skimmed Milk Powder (trace elements); and (e) National Metrology  Institute  of Japan  (NMIJ, Tsukuba,  Ibaraki, Japan),  NMIJ CRM 7502-a Trace Elements in White Rice Flour.

#### 2.2.2.  Archived proficiency test materials – food matrices

Archived PT samples, previously circulated by the US FDA as part of the Food Emergency Response Network (FERN) program’s assessment of laboratory performance, were available for this study and were analyzed using (a) the Zmax™ and Emax™ XRF analyzers (b) the HD Mobile® analyzer and (c) the two different Niton™ XRF analyzers: XL3t and XL5 Plus. At Wadsworth, these PT samples were originally analyzed blinded using FERN Method CHE.0009.02 (Sisko et al., 2024). In brief, aliquots of the  samples  were  digested  on  a  hot  block  and  then  analyzed  for multiple trace elements using an Agilent 7900 ICP-MS. The specific el- ements analyzed and reported differed for each PT sample. The specific FERN PT food rounds available included: summer sausage samples (n = 8)  circulated  in  Fiscal Year  (FY)  2022  (2022-FERN)  for  toxic metals analysis (Pb, Cd, and Tl); chicken cordon bleu samples (n = 8) circulated in FY2023 (2023-FERN) for toxic metals analysis (Pb, Tl, Hg, Cd, and Sn); and catfish fillet samples (n = 8) circulated in FY2024 (2024-FERN) for toxic metals analysis (W and Tl). These PT materials were prepared under the auspices of the Moffett Proficiency Testing Laboratory, for Food  and  Drug  Administration  (FDA)’s  Center  for  Food  Safety  and

Applied Nutrition  (CFSAN)  and  distributed to  US  state-based labora- tories for analysis. After participant laboratories had reported results to the  FERN,  a  final  performance  report  was  issued  that  included  the nominal target (spike) values, the assigned value based on ISO Algo- rithm   A,   standard   uncertainty,   and   z-scores   based   on   the   ISO 13528:2022 standard (ISO, 2022). Homogeneity and stability of these PT materials were established by the Moffett Laboratory according to the ISO/IEC 17043:2023 standard (ISO, 2023) and the statistical eval- uations based the ISO 13528:2022 standard (ISO, 2022).

#### 2.2.3.  Medicinal herbs and food samples

Various   medicinal   herbs   and   food   samples   were   procured   in Colombian and US stores based on their significance for health and food safety purposes. They included 1) Nine dried medicinal herb samples from Colombia: Cardamom (Elettaria cardamomum) (n = 1), Chamomile (Matricaria chamomilla) (n = 1), Eucalyptus (Eucalyptus spp.) (n = 1), Hibiscus Flower (Hibiscus sabdariffa) (n = 1), Holy Basil (Ocimum ten- uiflorum) (n = 3), Moringa (Moringa oleifera) (n = 1), and Pot marigold (Calendula officinalis) (n = 1); 2) Thirty dried spice samples from the Colombian and the US markets: Parsley (Petroselinum crispum) (n = 2), Achiote (fruit, Bixa orellana) (n = 1), Allspice (fruit, Pimenta dioica) (n = 1), Dreaming seasoning (spice blends) (n = 1), Black pepper (fruit, Piper nigrum) (n = 1), Chili Sakthi (Spice blends) (n = 1), Cinnamon (bark, Cinnamomum verum)  (n  =  1),  Clove  (Syzygium  aromaticum)  (n  =  1), Coriander (Coriandrum sativum) (n = 3), Curry (Murraya koenigii) (n = 1), Flavor Forward (Spice blends) (n = 1), Garam-Masala (Spice blends) (n = 1), Green tea (Camellia sinensis) (n = 1), Laurel (leaf, Laurus nobilis L.) (n = 1), Mural of flavor (herby blends) (n = 1), Nutmeg (Myristica fragrans)  (n  = 1),  Oregano  (Origanum vulgare)  (n  = 4),  Potato weed (Galinsoga parviflora) (n = 2), Red chili pepper (Capsicum annuum) (n = 1), Rosemary (Rosmarinus officinalis) (n = 1), Thyme (leaf, Thymus vul- garis) (n = 2), and Turmeric (Curcuma longa) (n = 1).

### 2.3.  Sample preparation

CRMs  were   analyzed   directly   on   the   Z-Spec  XRF   instruments (Zmax似 and Emax似) after transferring an aliquot of the sample to the XRF cup without any pretreatment. FERN PT food samples were sup- plied as homogeneous pastes, and an aliquot was transferred directly to the appropriate XRF sample cup for analysis on the Z-Spec, XOS HD Mobile®, and Niton似 XRF instruments. For the latter three, only the

FERN  2023  PT  samples,  which  were  spiked  at   μg/g  levels,  were analyzed.

Dried herb and spice samples were ground into a powder using a mortar and pestle and an aliquot transferred to the appropriate XRF sample cup for analysis on the Z-Spec XRF instruments. Some herb/spice samples that, upon an initial analysis, yielded unusual high results were re-ground to eliminate potential hot spots, and then reanalyzed on the Z- Spec XRF instruments.

For analysis on the Zmax似, approximately 150–170 mg of sample material was  transferred into  a  customized  sample  cup. The  Zmax似sample cup design comprises a plastic cylinder secured on one end with polypropylene  X-Ray  thin  film  of  12-μm  thickness  (Prolene®  film, Chemplex  Industries,  Inc.,  NY,  USA).  The  sample  (~150  mg)  was deposited onto the film, and carefully secured by a second Prolene® film, using light pressure to create a thin sample. The entire sample cup and the second film were secured by a plastic ring (Fig. 2a). The Zmax似sample cup is designed to achieve an ~0.2-mm sample thickness sand- wiched  between  the  two  film  layers  to  reduce  scattering  from  the sample.

In contrast, the sample cup arrangement for analysis on the Emax似utilizes plastic cup with a flat Prolene® X-Ray thin film window at one end. Approximately 0.5 – 1g of sample is transferred into the Emax似cup, which is secured with a plastic clip. For the Emax似 the energy of incident beam is much higher, such that a relatively thicker sample can be used to obtain a higher Cd fluorescence intensity. (Fig. 2b).

## 3.  Results and discussion

### 3.1.  Limits of detection for XRF analyzers

The LODs for the Zmax似 and Emax似, the HD Mobile® and the two Niton似 XRF analyzers are provided in Table 1. The LOD data for the Zmax似 and Emax似 were obtained based on the harmonized IUPAC protocol for a single laboratory method validation  (Thompson et al., 2002). This approach is based on repeated measurements of the analyte (s) at low levels in a relevant food matrix (n ≥ 6). The LOD is calculated as three times the standard deviation (SD). The data for the HD Mobile® were taken from the literature and represent LODs obtained in biological matrices using the plastic calibration mode (Guimar2016), and calculated  using  the  IUPAC  protocol.  The  LOD  data  for  the  Thermo Niton似 XL3t and XL 5 Plus were obtained in food matrices and based on

Fig. 2. Z-Spec XRF sample cup design and the sample ports for (a) the Zmax and (b) the Emax XRF analyzers.

Tabıe 1

Limits of detection (μg/g) reported for the E-max, Z-max, HD Mobile, Niton XL3t and Niton XL 5 Plus XRF instruments.

|  | Emax | Zmax | HD Mobile | Niton XL3t | Niton 5+ |
| --- | --- | --- | --- | --- | --- |
| As |  | 0.03 | 1 |  |  |
| Br | 6 | 2 |  |  |  |
| Ca | 240 | 15 |  |  |  |
| Cd | 0.04 |  | 11 | 6 | 5 |
| Cl |  | 200 |  |  |  |
| Cu | 2.5 | 0.4 | 1 | 4 | 3 |
| Fe | 17 | 0.9 | 6 | 9 | 9 |
| Hg | 0.15 | 0.06 | 1 | 2 | 1 |
| K | 1300 | 20 |  |  |  |
| Mn | 8 | 0.3 | 2 |  | 14 |
| Ni |  | 0.2 | 9 | 6 | 2 |
| Pb |  | 0.06 | 1 | 2 | 1 |
| Rb | 0.6 | 2 |  |  |  |
| Se | 1.3 | 0.07 | 1 | 2 | 1 |
| Sn | 0.11 |  |  | 10 | 7 |
| Sr | 1 | 0.9 | 2 |  |  |
| Tl |  | 0.06 |  |  |  |
| W | 0.2 | 0.1 |  |  |  |
| Zn | 10 | 0.1 | 2 | 5 |  |

the IUPAC protocol described above.

### 3.2.  Analysis of certified reference materials

The CRMs were analyzed in duplicate to assess trueness (accuracy), but only where certified values and traceability were available for ele- ments detectable by XRF (Table 2). Fig. 3 shows difference plots (by element) generated for the Z-Spec XRF reported values minus the CRM assigned values (y) as a function of mass fraction (x) i.e., the assigned CRM value for each element measured. The data in Fig. 3 show good agreement between the XRF found values and CRM assigned values for the 8 elements studied (Pb, Cd, Hg, As, Mn, Ni, Cu and Zn) and no ev- idence of any systematic bias. For Hg, there is a single data point with a small negative bias at the highest concentration measured, while for Cu, As and Ni, there is a single data point with a small positive bias at the highest concentration. However, when these data are expressed as the % difference relative to the CRM assigned value they are within ±10 %, which is remarkably good. The % difference plots are shown in Sup- plementary Fig. S1. In general, these data represent a robust validation of the Z-Spec FP algorithms optimized for the Zmax™ and Emax™ XRF analyzers, at least for those elements reported here and in ideal CRM matrices.

### 3.3.  Analysis of archived FERN PT samples

The  PT  food  samples  were  originally  provided  by  FDA  to  FERN laboratories across the US as part of a required performance assessment using FDA-approved methods for food matrices. Results were reported to FERN and scored against the assigned value by the scheme organizer. The  laboratory  performance  achieved  using  ICP-MS were within  the tolerance ranges established by the scheme organizer. The remaining PT sample  materials were  archived  for  future  method  development  and validation studies and were available for evaluating the Z-Spec M-XRF analyzers. The FERN PT samples (in the same sample cups used for the Emax™) were also re-analyzed on the HD Mobile®, as well as on the two Thermo Niton™ XRF analyzers, for comparison purposes. The data are reported in Tables 3–5 based on the FERN FY distribution. The specific archived  FERN  PT  samples  that  were  available  for  the  study  were labelled:  FERN-2022  (summer  sausage),  FERN-2023  (chicken  cordon bleu), and FERN-2024 (raw catfish fillet), consistent with the US fiscal year (FY) in which the PT study was conducted.

#### 3.3.1.  FERN-2022 PT samples: summer sausage

The FERN-2022 PT samples were a homogenized summer sausage matrix comprising: (a) two control samples (b) three samples supple- mented (i.e., spiked); with 150 μg/kg of Pb, Cd, and Tl; and (c) three samples supplemented with 1000 μg/kg of Pb and Tl (Table 3). As ex- pected, the Z-Spec XRF analyzers reported ζnon-detects’, i.e., below the LOD  for  Pb,  Cd,  and  Tl  in  the“un-spiked”control  samples.  For  the samples“nominally”spiked  with  150  μg/kg,  XRF  recoveries  ranged from 98.1 % to 109 % for the three metals, while precision (defined as the coefficient of variation, i.e., %CV) ranged from 7.0 % to 11.7 %. For the samples“nominally”spiked with 1000 μg/kg Pb and Tl, recoveries ranged from 98.1 % to 104 %, and the CV ranged from 1.02 % to 1.41 %. The FERN-2022 PT samples were not analyzed on either the HD Mo- bile® or Niton™ XRF analyzers.

#### 3.3.2.  FERN-2023 PT samples: chicken cordon bleu

The FERN-2023 PT sample set were a homogenized chicken cordon bleu  matrix  comprising:  (a)  two  control  samples;  (b)  three  samples spiked with  150  μg/kg Pb, Tl, Hg, Cd, and  1000  μg/kg Sn;  (c) three samples spiked with  1000  μg/kg Pb, Tl, Hg, Cd, and  2500  μg/kg Sn (Table  4).  Due  to  the high levels  of Sn  supplemented  into  the FY23 samples, they were selected for comparing the two Z-Spec detectors with the HD Mobile® and Niton™ analyzers. Those data are provided in the supplementary material (Table S1). The default analysis routines for the HD  Mobile®  and  Niton™  detectors  do  not  include  Tl  in  the  sample output and therefore those data are unavailable for comparison. For all analyzers except the HD Mobile®, the two control samples registered as non-detects for all elements of interest. The HD Mobile® registered a positive detection in the control samples for 26300 μg/kg for Sn. For the samples nominally spiked at 150 μg/kg, recoveries ranged from 90.7 % to 128 % for the Z-spec analyzers, while the repeatability ranged from 4.02 % to  13.5 %. For the Niton™  analyzers, all samples were non- detects for the elements of interest while for the HD Mobile®, Sn was recorded at 27200 μg/kg with other elements being non-detects. For the samples spiked at 1000 μg/kg, recoveries ranged from 88.2 % to 123 % for the Z-Spec analyzers, and repeatability (CV) ranged from 1.16 % to 12.1 %; for Sn, which was spiked at 2500 μg/kg, recovery was 97 % with 1.82 % %CV. For the Niton™ analyzers all elements except Sn were non- detects with the Sn recovery of 141.78 % and 65.31 % with 1.487 and 7.069 % CV for the Niton™ XL 5 Plus and the Niton™ XL3t respectively. For the HD Mobile® analyzer, Cd was not detectable; Pb had a recovery of 122.2 % with a 6.6 %CV, Hg had a recovery of 212.2 % with a 5.95 %

CV. As with previous measurements with the HD Mobile® Sn was very high at 62,500 μg/kg.

#### 3.3.3.  FERN-2024 PT samples: raw catfish fillet

The FERN-2024 sample set comprised homogenized raw catfish fillet supplemented with W and Tl. The set included: (a) two control samples; (b) three samples spiked with 150 μg/kg Tl and 1000 μg/kg W; and (c) three samples spiked with 500 μg/kg W and 1000 μg/kg Tl (Table 5). Recoveries for the Z-Spec analyzers ranged from 97.7 % to 116 %, and repeatability from 0.122 % to 6.88 %.

Overall recoveries for the elements detected (Pb, Cd, Sn, Tl, Hg, and W) were within ±20 % of the assigned value, with repeatability better than  10 % (CV) for the Z-Spec detectors, demonstrating performance that was judged fit for purpose. Moreover, XRF results generated by the Zmax™ and Emax™ devices were all found to be within the acceptable ranges  established  by  FDA  for  FERN  laboratories  using   approved methods based on ICP-MS. These data show a substantial improvement over previous XRF analyzers and detection capability as seen with the comparison  of  the  2023  PT  samples  where  the  older  XRF  devices struggled with detection or exhibited large differences for all measured samples. As with all analytical methods, repeatability (%CV) becomes inflated as concentrations approach the method LOD, yet the Zmax™ and  Emax™  devices  appear  suitable  for  the  rapid  screening  of food matrices for the toxic elements reported here. It is important to point out

B. Johnson-Restrepo et al.                                                                                                                                                                                        Radiation Physics and Chemistry 234 (2025) 112749

Tabıe 2

Multielement analysis of Certified Reference Materials by Z-Spec M-XRF Analyzer.

CRM Assigned valuea  (μg/g)                                 Z-Spec XRFb  (μg/g)                          CRM Assigned valuea  (μg/g)                           Z-Spec XRFb  (μg/g)                          CRM Assigned valuea  (μg/g)                        Z-Spec XRFb  (μg/g)

mean                %CV        %rec                                                                                           mean                %CV         %rec                                                                                        mean                %CV        %

rec

Cl

K

Ca

Mn

Fe

Ni

Cu

Zn

As

Se

Br

Rb Sr Hg Pb

Cl

K

Ca

Mn

Fe

6

Ni

Cu

Zn

As

Se

Br

Rb Sr Hg Pb

Cl

K

Ca

Mn

Fe

Ni

Cu

Zn

As

Se

Br

Rb Sr Hg Pb

Cl

K

NMIJ 7502 (white Rice)

ERM BD151 (skimmed milk powder)

NIST SRM 1515 (apple leaves)

NIST SRM 1577c (Bovine Liver)

1430  60.0 11.2

4.48	0.39	3.02 26.0

0.109

0.068

0.0043

9.8 17

13.9 0.29

53

5.0 44.9

0.19

0.52 0.207

582

16080

15250  54.1 82.7

0.94 5.69

12.45

10.2 25.1

0.47

10230

115

1190

55

11

3.9

0.37

3.03 25.1

0.09 0.01 0.35 1.68

ND

ND

10.4 16.8 14.2	0.3 50.2	0.1		5.1 44.8

0.2 12.7 17.8 		3.6 	0.44	0.3

642

16891

16964

55

77

1.00		5.88 13.50	0.02	0.04			1.69	9.96

25.05

0.46

3092

11255

83.2	91.7	94.2	87.5	94.9 100.2

86.2

3.09 0.18 0.00 0.40 1.08 7.64 2.10 0.59 7.52 0.00 1.00

0.21 ND

ND

106

99

102

107

95

102

100

97

84

126

110

105

111

102

93

106

103

108

1.14 1.11 1.14 0.00 0.38 5.24 0.28 0.27 3.84 0.21 0.20 0.45 1.46 2.72

98

100

97

0.66	0.09	0.07	0.27	0.18	0.71	0.12	0.00	3.29 16.84	0.29	0.02	0.20	2.48

110

2.54 2.10

NIST SRM 1573a (tomato leaves)

NIST SRM 1570a (spinach leaves)

IAEA-413 (Algae)

BOVM_1 (Bovine Muscle)

6600

26760

50450  246.3 367.5

1.58 	4.70 30.94 	0.113

0.054

14.8 85

29000

15260

76

2.14 12.22 82.3

0.068

0.115

55.54

10740

3143

158

1370

113

11.1 169  127

53.2 242

1900.0 15200

6895

25820

48590

236

330

1.6	4.7 31.4	0.1

ND

1278

26

84

7064

27533

14892

76

252

2.07 12.1

79.3

0.06 	0.123 35.75 13.30 53.81 	0.05

0.21

708

11037

3295

169

1376

114

7.08 166

128

3.51		3.15	8.95 47.6

242

2236

16472

104  	96.5	96.3	95.8	89.9	98.0 101  102  		78.2

174  99.0

1.63		1.54		1.90		1.48		1.21	3.65		1.49	0.65 11.25 ND

0.47 0.83 0.49

95

98

100

97

99

96

94

106

97

1.62		1.31		1.58	0.46	0.87	0.68	0.59	0.31	4.42 12.12	0.08	0.06	0.14 18.86 13.40

103

105

107

100

101

64

98

100

89

100

1.20 0.68 0.67 0.10 0.21 0.15 0.80 0.28 0.29 0.81 0.16 0.09 0.01

0.08

108

1.52 0.84

NIST SRM 1547 (peach leaves)

NIST SRM 1568b (rice flour)

NIST SRM 2976

(Mussel Tissue)

NIST SRM 2977

(Mussel Tissue)

111

101

105

101

96

134

109

106

100

99

102

95

87

95

111

103

99

107

99

96

87

91

103

106

105

94

89

111

98

119

107

101

101

100

176

73

93

105

361

24330

15590

1.41		1.14		1.31	0.62	0.59	3.07	0.86	0.48 ND

2.96 0.24 0.08 0.10

2.83

0.00	0.06	0.00	0.07	0.10	7.44	0.85	0.00	4.63		1.33	0.32	0.38	0.39 24.96

0.02	0.25	0.17	0.10	0.22	0.00	0.66	0.19	0.21	0.94	0.04	0.96	0.06 13.8 		1.87

401

24454

16326

97.8 219.8

98.5

210

0.689		3.75 17.97 	0.06

0.12

19.7 53.0

0.92	4.1 19.0

ND

0.87

0.12 11.4 19.4 54.2

0.83

301.1 1282

118.4 19.2

261

1214

132

7.42

2.35 19.4

0.285 0.37 8.31 6.20

0.008

19.7

7.33	0.19	2.51 19.26	0.28	0.32	7.53	6.35	0.18

0.009

57000

9700

7600

33

171

0.93

4.02 137

13.3 1.8

4.14 93.0

0.061

1.19

59865

9094

6788

37

167

1.1 4.3

138.38 13.4 	1.8

288.28 7.3

68.3

0.057 1.25

45640                  1.03

12374                1.26

(continued on next page)

Tabıe 2 (continued )

a   CRM assigned values (certified, reference, or other) taken from the NMI certificate of analysis.

b   Z-            -SpecE max used to report Cd and Sn, and the Z-Spec Z-max used to report all others.

| % rec |  | 102 | 95 | 100 | 104 | 101 | 97 | 102 |  |  | 97 |  | 96 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Z-Spec XRF(μg/g)b mean %CV | 8895 1.36 | 24.4 0.15 | 261 0.60 | 6.08 0.58 | 9.8 0.07 | 137 0.48 | 8.61 1.66 | 1.82 0.19 | 229 0.29 | 10.8 2.36 | 67.4 0.17 | 0.10 10.21 | 2.17 11.34 |
| CRM Assigned valuea (μg/g) |  | 23.9 | 274 | 6.06 | 9.42 | 135 | 8.83 | 1.78 |  |  | 69.3 |  | 2.27 |
| %rec | 120 | 114 | 99 | 470 | 111 | 100 |  | 101 | 90 | 100 | 265 |  | 78 |
| Z-Spec XRFb (μg/g) mean %CV | 175 1.22 | 0.42 6.73 | 70.5 0.52 | 0.24 3.01 | 3.15 1.12 | 142 0.15 | ND ND | 0.077 3.67 | 0.995 1.85 | 28.7 0.11 | 0.14 2.05 |  | 0.30 10.00 |
|  | 145 | 0.37 | 71.20 | 0.05 | 2.84 | 142 | 0.009 | 0.076 | 1.1 | 28.7 | 0.052 |  | 0.38 |
| CRM Assigned valuea (μg/g) |  |  |  |  |  |  |  |  |  |  |  |  |  |
| %rec | 105 | 109 | 105 | 135 |  | 102 | 107 | 97 |  |  | 147 |  | 104 |
| Z-Spec XRFb (μg/g) mean %CV | 137 2.06 | 11.4 0.68 | 207 0.94 | 0.060 47.14 | 277 0.33 | 185 0.37 | 0.02 13.47 | 2.0 0.18 | 11.9 0.26 | 34.2 0.18 | 0.14 0.50 |  | 0.065 30.46 |
|  | 131 | 10.46 | 198 | 0.045 |  | 181.1 | 0.0196 | 2.03 |  |  | 0.0953 |  | 0.063 |
| CRM Assigned valuea (μg/g) |  |  |  |  |  |  |  |  |  |  |  |  |  |
|  | Ca | Mn | Fe | Ni | Cu | Zn | As | Se | Br | Rb | Sr | Hg | Pb |

that, while the XRF analyst was“blinded”to the assigned values for	these archived PT samples, it was not possible to analyze them under “true”PT conditions.

### 3.4.  Analysis of medicinal herb samples

A variety of medicinal herb samples were procured from Colombian and US based stores and analyzed in duplicate for a suite of essential and non-essential elements using the Zmax™ and Emax™ XRF analyzers. Samples analyzed included: a) Cardamom, b) Chamomile, c) Eucalyptus, d) Hibiscus flower, e) Holy basil, f) Moringa, and g) Pot marigold. Lead and Cd were detected in most samples, with Pb levels ranging from non- detect (moringa) to 0.96 μg/g (holy basil), and Cd levels ranging from non-detect to 0.31  μg/g  (holy basil). Two samples of holy basil con- tained As at low levels, 0.10 and 0.52 μg/g. Neither Hg nor Sn were detected in  any of the herbal samples. The complete table of results (mean, n = 2) of the analysis of medicinal herb samples is reported in the supplementary material. (Table S2).

In recent years, the popularity of herbal products has surged, with many consumers seeking alternatives to conventional pharmaceutical treatments. However, it is important to highlight that medicinal herbs are  often  sold  without  undergoing  the  necessary  testing  to  ensure compliance with good agricultural and manufacturing practices (GAP and GMP). This lack of regulation and oversight can pose significant risks to public health, as it allows for the possibility of contamination with harmful substances, including toxic heavy metals and pesticides (Jordan et al., 2010; Ruparel and Lockwood, 2011; Zhang et al., 2012). One  of the key  concerns  regarding medicinal herb  is  the  absence  of standardized testing  to verify  that these products  are  free  from  con- taminants. Unlike pharmaceuticals, which undergo rigorous testing for safety,  efficacy,  and  quality  control,  herbal  products  are  subject  to minimal regulation in many parts of the world. As a result, manufac- turers may not consistently follow best practices for cultivation, har- vesting, or processing, leading to potential issues related to the safety and quality of the final product. The lack of testing for contaminants such as heavy metals further exacerbate this issue, as consumers may unknowingly be exposed to harmful substances that can have serious long-term health effects (Kosalec et al., 2009). Heavy metals, such as Pb, As, Hg, and Cd, are a particular concern due to their toxicity even at low levels  of exposure.  Improper  storage, handling,  or processing  of me- dicinal herbs can also introduce toxic metals into the final product.

### 3.5.  Analysis of the dried spice samples

A limited number (n = 22) of dried spice samples (rosemary, thyme, clove,  oregano,  cinnamon,  curry,  and  others),  were  procured  from Colombia and US based stores, and analyzed in duplicate for a suite of essential and non-essential elements using the Emax™ and Zmax™ M- XRF  analyzers.  The  complete  table  of results  (mean,  n  =  2)  of the analysis  of spice  samples  is  reported  in  the  supplementary  material. (Table 2S). Overall, essential element content decreased in the following order: Fe > Mn > Cu > Zn > Ni > Se. Although Cd and As were detected in some samples, the highest concentrations of As were found in two brands of oregano at 2.44 and 2.98 μg/g, while the highest Cd levels were found in two brands of potato weed at 0.99 and 1.04 μg/g. Rela- tively high concentrations of Pb were found in two different brands of thyme, at 19.7 and 16.3 μg/g, as well as in oregano and coriander at levels of 7.17 and  7.87  μg/g, respectively. The lowest amount of Pb detected was found in chili spice blend at 0.09 μg/g. Mercury was not detected in any of the spice samples.

The metal  content  of spices  is  influenced by  a variety  of factors, including the species of plant, the specific plant regions harvested, as well as local environmental conditions such as soil metal content, pH, water contamination, air pollution, industrial activities, and the type of fertilizers used (Tokal1o˘glu et al., 2018). Trace elements like Pb, Cd, As, and Hg can accumulate in many plant species through the process of

Fig. 3. Difference  plots  by  element  showing  the  Z-Spec  XRF  measured  values  minus  the  CRM  assigned  values.  All  elements  except  Cd  measured  using  the Zmax™ XRF.

uptake from the soil. In general, the root system of most plants absorbs the highest levels of toxic metals because it is in direct contact with the soil. This makes the roots a primary site for metal uptake (Shahid et al., 2017).

## 4.  Conclusions

Second generation M-XRF instruments (Zmax™ and Emax™) were evaluated for accuracy and repeatability. In addition to analyzing food- based CRMs, a selection of archived food PT samples provided by the FDA for FERN laboratories were available to judge M-XRF performance. Overall, the Zmax™ and Emax™ performance for the elements detected was within ±20 % of the assigned values and met the FDA’s PT per- formance criteria for FERN laboratories using an FDA-approved method (ICP-MS).  Based  on  these  data,  the  Zmax™  and  Emax™  M-XRF  in- struments were judged fit for purpose for the elements detected. The

LODs for the Zmax™ and Emax™ M-XRF instruments were up to 100 times lower for a suite of toxic elements (e.g., Pb, Cd, Hg, As, Tl, W, and Sn) compared to those obtained previously for the XOS HD Mobile® M- XRF  system and two conventional Niton™ XRF analyzers. The LODs were consistent with performance observed for multiple CRMs as well as for  several  archived  FERN  PT  materials.  The  two  Z-Spec  M-XRF  in- struments were also used to screen a range of“real-world”medicinal herbs  and  spices  sourced  from  US  and  Colombian  stores.  While  the Zmax™  and Emax™  instruments  demonstrated reasonable LODs,  ac- curacy and repeatability for the key toxic elements (As, Cd, Hg and Pb), optimal performance was achieved by selecting the Emax™ for Cd (and Sn), and the Zmax™ for As, Hg, and Pb.

## CRediT authorship contribution statement

Boris  Johnson-Restrepo:  Writing   –   original   draft,  Validation,

Table 3

Analysis of FERN FY2022 PT samples (summer sausage) for toxic metals using the Z-Spec M-XRF analyzersb.

| FERN PT ID | Element | Spike Level (μg/ kg) | FERN Assigned value (μg/ kg) | SE (μg/ kg) | Acceptable range (μg/ kg) | XRF meana (μg/ kg) | SE (μg/ kg) | CV (%) | Recovery (%) |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Blank | Pb Cd Tl |  |  |  |  | ND ND ND |  |  |  |
| C-03 & | Pb | Un-spiked |  |  |  | ND |  |  |  |
| C-05 | Cd | Un-spiked |  |  |  | ND |  |  |  |
|  | Tl | Un-spiked |  |  |  | ND |  |  |  |
| C-01, C-04 & C- | Pb | 150 | 160 | 2 | 138–182 | 155 | 6.0 | 11.7 | 103 |
| 07 | Cd Tl | 150 150 | 157 148 | 3 2 | 123–191 124–172 | 151 164 | 3.7 3.2 | 7.3 7.0 | 100 109 |
| C-02, C-06 & C- | Pb | 1000 | 1034 | 12 | 904–1164 | 1038 | 3.5 | 1.0 | 104 |
| 08 | Cd Tl | Un-spiked 1000 | 998 | 14 | 846–1150 | ND 981 | 4.6 | 1.4 | 98 |

a  n = 3.

b  XRF analysis for Cd by E-max, Pb and Tl by Z-max.

Table 4

Analysis of FERN FY2023 PT samples (chicken cordon bleu) for toxic metals using the Z-Spec M-XRF analyzersb.

| FERN PT ID | Element | Spike Level (μg/ kg) | FERN Assigned value (μg/ kg) | SE (μg/ kg) | Acceptable range (μg/ kg) | XRF meana (μg/ kg) | SE (μg/ kg) | CV (%) | Recovery (%) |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Blank | Pb |  |  |  |  | ND | ND |  |  |
|  | Tl |  |  |  |  | ND | ND |  |  |
|  | Hg |  |  |  |  | ND | ND |  |  |
|  | Cd |  |  |  |  | ND | ND |  |  |
|  | Sn |  |  |  |  | ND | ND |  |  |
| C-01 & C-03 | Pb | Un-spiked |  |  |  | ND | ND |  |  |
|  | Tl | Un-spiked |  |  |  | ND | ND |  |  |
|  | Hg | Un-spiked |  |  |  | ND | ND |  |  |
|  | Cd | Un-spiked |  |  |  | ND | ND |  |  |
|  | Sn | Un-spiked |  |  |  | ND | ND |  |  |
| C-02, C-04 & C- 06 | Pb | 150 | 157.7 | 2.42 | 132.9–182.5 | 163 | 3.5 | 6.4 | 109 |
|  | Tl | 150 | 152.1 | 1.71 | 134.0–170.1 | 136 | 6.1 | 13.5 | 91 |
|  | Hg | 150 | 154.8 | 1.6 | 137.3–172.3 | 191 | 2.6 | 4.0 | 128 |
|  | Cd | 150 | 158.6 | 1.31 | 142.6–174.6 | 146 | 4.6 | 9.5 | 98 |
|  | Sn | 1000 | 1101.3 | 31.2 | 891.7–1310.9 | 1065 | 42.8 | 12.1 | 107 |
| C-05, C-07 & C- 08 | Pb | 1000 | 989.6 | 8.98 | 890.6–1088.7 | 1044 | 4.7 | 1.3 | 104 |
|  | Tl | 1000 | 980 | 12 | 852.7–1107.3 | 882 | 3.4 | 1.2 | 88 |
|  | Hg | 1000 | 990.3 | 9.42 | 875.6–1105.0 | 1230 | 7.5 | 1.8 | 123 |
|  | Cd | 1000 | 1004.1 | 7.67 | 916.2–1091.9 | 1083 | 9.9 | 2.7 | 108 |
|  | Sn | 25000 | 25325.9 | 775.7 | 19881.2–30770.6 | 24239 | 147 | 1.8 | 97 |

a  n = 3.

b  XRF analysis for Cd by E-max, Pb and Tl by Z-max.

Table 5

Analysis of FERN FY2024 PT samples (raw catfish filet) for toxic metals using the Z-Spec M-XRF analyzersb.

| FERN PT ID | Element | Spike Level (μg/ kg) | FERN Assigned value (μg/ kg) | SE (μg/ kg) | Acceptable range (μg/ kg) | XRF meana (μg/ kg) | SE (μg/ kg) | CV (%) | Recovery (%) |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Blank | W |  |  |  |  | ND |  |  |  |
|  | Tl |  |  |  |  | ND |  |  |  |
| C-05 & C-06 | W | Un-spiked |  |  |  | ND |  |  |  |
|  | Tl | Un-spiked |  |  |  | ND |  |  |  |
| C-01, C-04 & C- | W | 1000 | 1123 | 48.6 | 744.1–1501 | 1029 | 0.1 | 0.5 | 103 |
| 07 | Tl | 150 | 153 | 1.7 | 135.1–170.4 | 163 | 0.1 | 4.5 | 109 |
| C-02, C-03 & C- | W | 500 | 554 | 23.2 | 371.4–736.5 | 579 | 0.2 | 0.1 | 116 |
| 08 | Tl | 1000 | 988.5 | 1.0 | 880.1–1096.2 | 977 | 0.2 | 6.9 | 98 |

a  n = 3.

b  XRF analysis for Cd by E-max, Pb and Tl by Z-max.

Methodology, Investigation, Formal analysis, Data curation, Conceptu- alization. Ezekiel Blain: Writing – review & editing, Validation, Meth- odology,  Investigation,  Formal  analysis,  Data  curation.  Christopher Judd: Writing – review & editing, Methodology, Investigation, Formal analysis, Data curation. Andrew Tysoe: Methodology, Formal analysis. Patrick J. Parsons: Writing – review & editing, Writing – original draft,

Validation, Supervision, Project administration, Methodology, Investi- gation, Formal analysis, Data curation, Conceptualization.

## Declaration of competing interest

The authors declare that they have no known competing financial

interests or personal relationships that could have appeared to influence the work reported in this paper.

## Acknowledgments

The authors thank Z-Spec Inc. (East Greenbush, NY) for the loan of the Zmax™ and Emax™ M-XRF instruments for evaluation purposes. We are grateful to Dr. Zewu Chen, Z-Spec Inc., for technical assistance and helpful discussions. This work was supported in part by the Food and Drug Administration (FDA) of the U.S. Department of Health and Human Services (HHS) as part of a financial assistance award U19FD007089 (Parsons). We are grateful to the US FDA for permission to use data from the FERN PT reports as part of this study. The contents are those of the author(s) and do not necessarily represent the official views of, nor an endorsement, by FDA/HHS, or the U.S. Government. Use of trade names is for identification purposes only and does not imply an endorsement by the New York State Department of Health. Dr. Boris Johnson-Restrepo acknowledges support from the Ministry of Science, Technology, and Innovation (MinCiencias) and Sistema General de Regalías de Colombia under grant BPIN- 2020000100427 (Gobernaciones de Atl´antico, Bolí- var, Cesar, C´ordoba, La Guajira, Magdalena, Sucre, and San Andr´es y Providence, Colombia).

## Appendix A.  Supplementary data

Supplementary data to this article can be found online at https://doi. org/10.1016/j.radphyschem.2025.112749.

## Data availability

Data will be made available on request.

Chen, Z.W., Wittry, D.B., 1998. Microanalysis by monochromatic microprobe x-ray  fluorescence - physical basis, properties, and future prospects. J. Appl. Phys. 84, 1064–1073.

de Boer, D.K.G., 1989. Fundamental parameters for X-ray fluorescence analysis. Spectrochim. Acta B Atom Spectrosc. 44, 1171–1190.

Fleming, D.E.B., Crook, S.L., Evans, C.T., Nader, M.N., Atia, M., Hicks, J.M.T.,

Sweeney, E., McFarlane, C.R., Kim, J.S., Keltie, E., Adisesh, A., 2020. Portable X-ray fluorescence of zinc applied to human toenail clippings. J. Trace Elem. Med. Biol. 62, 126603.

Fleming, D.E.B., Nader, M.N., Foran, K.A., Groskopf, C., Reno, M.C., Ware, C.S.,

Tehrani, M., , D., Parsons, P.J., 2017. Assessing arsenic and selenium in a single nail clipping using portable X-ray fluorescence. Appl. Radiat. Isot. 120, 1–6.

, D., Praamsma, M.L., Parsons, P.J., 2016. Evaluation of a new optic-enabled portable XRF instrument for measuring toxic metals/metalloids in consumer goods and cultural products. Spectrochim. Acta Part B At. Spectrosc. 122, 192–202.

ISO, 2022. Statistical Methods for Use in Proficiency Testing by Interlaboratory Comparison. ISO/TC 69/SC 6, Switzerland, p. 93.

ISO, 2023. Conformity Assessment — General Requirements for the Competence of Proficiency Testing Providers, p. 36. Switzerland.

Jordan, S.A., Cunningham, D.G., Marles, R.J., 2010. Assessment of herbal medicinal products: challenges, and opportunities to increase the knowledge base for safety assessment. Toxicol. Appl. Pharmacol. 243, 198–216.

Kalnicky, D.J., Singhvi, R., 2001. Field portable XRF analysis of environmental samples. J. Hazard Mater. 83, 93–122.

Kitov, B.I., 2000. Calculation features of the fundamental parameters in XRF. X Ray Spectrom. 29, 6.

Kosalec, I., Cvek, J., Tomi´c, S., 2009. Contaminants of medicinal herbs and herbal products. Arh. Hig. Rada. Toksikol. 60, 485–501.

Marguí, E., Queralt, I., de Almeida, E., 2022. X-ray fluorescence spectrometry for

environmental analysis: basic principles, instrumentation, applications and recent trends. Chemosphere 303, 135006.

McIntosh, K.G., , D., Cusack, M.J., Vershinin, A., Chen, Z.W., Yang, K.,

Parsons, P.J., 2016. Evaluation of portable XRF instrumentation for assessing

potential environmental exposure to toxic elements. Int. J. Environ. Anal. Chem. 96, 15–37.

Revenko, A.G., Pashkova, G.V., 2023. X-ray fluorescence spectrometry: current status and prospects of development. J. Anal. Chem. 78, 1452–1468.

Ruparel, P., Lockwood, B., 2011. The quality of commercially available herbal products. Nat. Prod. Commun. 6, 733–744.

## References

Acheampong, A., Sarpong, K., Akoto, O., Apau, J., Darko, G., 2024. Health risk

assessments of potentially toxic elements in selected medicinal plants from the ashanti region of Ghana. Chemistry Africa 7, 4361–4379.

Ahsan, H., 2011. Arsenic in drinking water. Kaohsiung J Med Sci 27 (9), 358–359. https://doi.org/10.1016/j.kjms.2011.05.002. PubMed PMID: 219.

Akhbarizadeh, R., Dobaradaran, S., Spitz, J., Mohammadi, A., Tekle-R¨ottering, A., De-la- Torre, G.E., Keshtkar, M., 2023. Metal(loid)s in herbal medicines and their infusions: levels, transfer rate, and potential risks to human health. Hygiene Environ. Health Adv. 5, 100042.

Chen, Z., Gibson, W.M., 2002. Doubly curved crystal (DCC) X-ray optics and applications. Powder Diffr. 17, 99–103.

Chen, Z.W., Gibson, W.M., Huang, H., 2008. High definition X-ray fluorescence: principles and techniques. X-Ray Optics and Instrumentation 2008, 318171.

Shahid, M., Dumat, C., Khalid, S., Schreck, E., Xiong, T., Niazi, N.K., 2017. Foliar heavy metal uptake, toxicity and detoxification in plants: a comparison of foliar and root metal uptake. J. Hazard Mater. 325, 36–58.

Singh, V.K., Sharma, N., Singh, V.K., 2022. Application of X-ray fluorescence

spectrometry in plant science: solutions, threats, and opportunities. X Ray Spectrom. 51, 304–327.

Sisko, P., Farrow, M., Gray, P., Mindak, W., Cheng, J., 2024. In: FERN), F.E.R.N. (Ed.), Inductively Coupled Plasma-Mass Spectrometric Determination of Arsenic,

Cadmium, Chromium, Lead, Mercury, and Other Elements in Food Using Microwave Assisted Digestion, vol. 38.

Thompson, M., Ellison, S.L.R., Wood, R., 2002. Harmonized guidelines for single-

laboratory validation of methods of analysis (IUPAC Technical Report). Pure Appl. Chem. 74, 835–855.

Tokalıo˘glu, ¸S ., Çiçek, B., ˙Inanç, N., Zararsız, G., ¨Oztürk, A., 2018. Multivariate statistical

analysis of data and ICP-MS determination of heavy metals in different brands of spices consumed in kayseri, Turkey. Food Anal. Methods 11, 2407–2418.

WHO, 2007. WHO Guidelines for Assessing Quality of Herbal Medicines with Reference to Contaminants and Residues. Geneva.

Zhang, J., Wider, B., Shang, H., Li, X., Ernst, E., 2012. Quality of herbal medicines: challenges and solutions. Compl. Ther. Med. 20, 100–106.